# 1. Scope & Planning

STEP 1 - Scope & Planning  |  CNC Machining - Ti-6Al-4V Femoral Stem

| Field | Detail |
| --- | --- |
| Process Name | CNC Machining - Ti-6Al-4V Femoral Stem |
| Scope Start | Raw Ti-6Al-4V Bar Stock Receipt |
| Scope End | Final Dimensional Inspection / Release |
| Device Classification | Class III - PMA (21 CFR 814) |
| Standard Applied | AIAG-VDA FMEA 4th Edition 2019 |
| Regulatory Framework | 21 CFR 820 (QSR) ¦ ISO 13485:2016 |
| Device Description | Cementless Ti-6Al-4V Femoral Stem - Implantable Hip Prosthesis |
| Team Members | Quality Engineer, Process Engineer, Supplier Quality, R&D |
| FMEA Facilitator | [Your Name] - Quality Engineer |
| FMEA Number | PFMEA-HIP-001 |
| Revision | Rev A |
| Date | 2026-06-22 |
| Next Review Date | 6 months post-transfer validation |
| Status | Draft - Internal Review |

AIAG-VDA FMEA 2019  |  21 CFR 820  |  ISO 13485

# 2. Structure Analysis

STEP 2 - Structure Analysis  |  System Hierarchy Breakdown

| System Level | Element Name | Function Contribution |
| --- | --- | --- |
| Level 1 - System | Manufacturing Cell | CNC Machining production environment |
| Level 2 - Subsystem | CNC Turning Station | Primary material removal and geometry creation |
| Level 3 - Component | Cutting Tool (Insert) | Executes material removal; controls Ra and geometry |
| Level 3 - Component | Workholding Fixture | Maintains part datum and prevents movement |
| Level 3 - Component | Coolant Delivery System | Controls chip evacuation and thermal stability |
| Level 3 - Component | CNC Part Program | Defines toolpath, feeds, speeds, and tolerances |
| Level 3 - Component | Operator / Setup Tech | Executes setup, monitors process, responds to alarms |
| Level 4 - Interface | CMM Inspection Station | In-process and final dimensional verification |
| Level 4 - Interface | Electronic Traveler (DHR) | Enforces operation sequence and records results |

AIAG-VDA FMEA 2019  |  21 CFR 820  |  ISO 13485

# 3. Function Analysis

STEP 3 - Function Analysis  |  Intended Functions & Requirements

| Process Element | Intended Function | Requirement Satisfied |
| --- | --- | --- |
| Cutting Tool (Insert) | Remove material to achieve OD +/-0.010 mm and Ra <= 0.8 um on taper surface | Device drawing tolerance stack; osseointegration surface finish spec |
| Workholding Fixture | Maintain part concentricity <0.005 mm TIR during all turning operations | Geometric dimensioning: perpendicularity and runout callouts |
| Coolant Delivery System | Deliver coolant at >=50 PSI to maintain <60C tool-chip interface temperature | Ti-6Al-4V machining process spec; tool life and metallurgical integrity |
| CNC Part Program | Execute validated toolpath per ECO-approved revision; enforce feed/speed limits | Validated manufacturing process (IQ/OQ/PQ); 21 CFR 820.70(i) software control |
| Operator / Setup Tech | Load correct program revision, set datum offsets, perform first-article inspection | Work instruction WI-CNC-001; operator qualification records OQR-2024 |
| CMM Inspection Station | Verify OD, ID, taper angle, and length within print tolerance post-machining | Inspection plan IP-HIP-003; acceptance criteria per device drawing Rev G |
| Electronic Traveler (DHR) | Enforce operation sequence gate checks; capture lot/SN, operator, and tool data | 21 CFR 820.184 Device History Record; ISO 13485 7.5.3 traceability |

AIAG-VDA FMEA 2019  |  21 CFR 820  |  ISO 13485

# 4. Failure Analysis

STEP 4 - Failure Analysis  |  Failure Modes, Effects & Causes

| Process Step | Failure Mode | Effect on Patient / Downstream Process | Severity (1-10) | Root Cause | Current Prevention Control | Current Detection Control |
| --- | --- | --- | --- | --- | --- | --- |
| CNC Turning - OD | OD out of tolerance (oversize) | Implant-stem mismatch; potential intraoperative fracture or non-union - patient harm | 8 | Tool wear beyond limit; incorrect offset entry; thermal expansion of Ti | Insert change interval SOP; offset verification at setup | Post-op CMM inspection; SPC chart OD trend |
| CNC Turning - Surface | Surface roughness Ra > 0.8 um on taper | Reduced osseointegration; long-term implant loosening - re-operation risk | 9 | Worn or chipped insert; inadequate coolant flow; excessive feed rate | Profilometer check per lot; insert change per interval | Profilometer measurement per part; visual inspection |
| Workholding - Fixture | Fixture slip / part shift mid-cycle | Geometric error cascade (runout, perpendicularity OOS); scrap or undetected escape | 7 | Improper clamping torque; fixture wear; incorrect locating pin seating | Torque spec on setup sheet; periodic fixture calibration | First-article runout check; operator pre-run confirmation |
| Coolant Delivery | Coolant pressure drop / flow loss | Thermal damage to Ti substrate; metallurgical alteration; implant scrapped or escaped | 9 | Clogged filter; pump failure; hose leak; operator valve error | Coolant PM schedule; filter change interval | Visual observation only - no automated interlock (GAP) |
| CNC Program - Revision | Wrong program revision loaded (obsolete program) | Incorrect geometry machined; toolpath collision risk; patient dimension non-conformance | 8 | Paper traveler allows manual program selection; no enforced program lock | Program naming convention; operator training | Setup verification checklist; first-article measurement |
| Inspection - CMM Step | CMM inspection step skipped / bypassed | Non-conforming part released to sterilization and distribution - direct patient risk | 7 | Production schedule pressure; traveler signature forgery; inadequate gate control | Work instruction WI-CMM-002; supervisor sign-off | DHR review at final release - detection is end-of-line only (GAP) |

AIAG-VDA FMEA 2019  |  21 CFR 820  |  ISO 13485

# 5. Risk Rating

AIAG-VDA 2019 Action Priority (AP) -- NOT Traditional RPN

AIAG-VDA 2019 replaces Risk Priority Number (RPN = SxOxD) with Action Priority (AP: High / Medium / Low). Rationale for medical devices: a failure with Severity 9 that is NEVER detected (D=10) must be actioned immediately regardless of occurrence rate - RPN arithmetic can mask this by averaging scores. AP weights Detection gaps more heavily, reflecting that an undetected failure reaching a patient is catastrophic. The AP lookup table (not multiplication) drives priority.

AP Lookup: S9-10+D7-10->H | S9-10+D4-6+O>=4->H | S7-8+D7-10+O>=4->H | All others->M or L

STEP 5 - Risk Rating  |  AIAG-VDA Action Priority Table

| Failure Mode | S (Severity 1-10) | O (Occurrence 1-10) | D (Detection 1-10) | Action Priority | AP Rationale | Legacy RPN (Reference Only) |
| --- | --- | --- | --- | --- | --- | --- |
| OD out of tolerance (oversize) | 8 | 4 | 3 | L | S=7-8 + D<=3: detection adequate for severity level | 96 |
| Surface roughness Ra > 0.8 um | 9 | 3 | 2 | M | S>=9 + D<=3: strong detection mitigates critical severity | 54 |
| Fixture slip / part shift mid-cycle | 7 | 3 | 4 | L | S=7-8 + D=4-6 + O<4: controlled risk | 84 |
| Coolant pressure drop / flow loss | 9 | 3 | 5 | M | S>=9 + D=4-6 + O<4: rare critical but detection acceptable | 135 |
| Wrong program revision loaded | 8 | 2 | 3 | L | S=7-8 + D<=3: detection adequate for severity level | 48 |
| CMM inspection step skipped | 7 | 4 | 7 | H | S=7-8 + D>=7 + O>=4: high-severity, frequent, poorly detected | 196 |

AIAG-VDA FMEA 2019  |  21 CFR 820  |  ISO 13485

# 6. Optimization

STEP 6 - Optimization / Corrective Actions  |  Risk Reduction Plan

| Initial AP | Failure Mode | Action Description | Action Owner | Target Date | Revised S | Revised O | Revised D | Revised AP |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| H | Coolant pressure drop / flow loss | Install in-line coolant flow sensor with machine interlock: auto-feed hold + alarm if PSI < 45. Validate per IQ/OQ protocol IQ-COOL-001. Update PFMEA and Control Plan CP-HIP-001. | Process Engineer | 2026-Q3 | 9 | 2 | 2 | M |
| H | Wrong program revision loaded | Implement electronic traveler (eDHR) with barcode-locked program selection. Only approved ECO revision loads; prior revisions auto-archived. Validate per 21 CFR 820.70(i). Training completion rate 100% before go-live. | Quality Engineer + IT | 2026-Q4 | 8 | 2 | 1 | L |
| M | Surface roughness Ra > 0.8 um | Reduce insert change interval from 150 to 100 parts; implement profilometer check every 50 parts (sample basis). Update PM schedule and WI-CNC-001 Rev B. | Quality Engineer | 2026-Q3 | 9 | 2 | 1 | M |
| H | CMM inspection step skipped | Add mandatory CMM gate-check in eDHR workflow: part cannot advance until CMM pass record is uploaded. CMM data auto-linked to DHR lot record. Supervisor override requires Level III QE approval. | Quality Engineer | 2026-Q4 | 7 | 3 | 2 | L |

AIAG-VDA FMEA 2019  |  21 CFR 820  |  ISO 13485

# 7. Results Summary

STEP 7 - Results & Documentation Summary

pFMEA Metrics Summary (auto-computed)

| Total Failure Modes Analyzed | 6 |
| --- | --- |
| Initial High (H) Action Priority | 1 |
| High AP After Actions Implemented | 0 |
| Medium (M) AP - Initial | 2 |
| Medium (M) AP - After Actions | 2 |
| Low (L) AP - After Actions | 4 |
| Total Corrective Actions Assigned | 4 |
| Actions with Owner Assigned | 4 |
| Actions with Target Date | 4 |

Linked Quality System Documents

| Document Type | Document Number | Title / Description | Status |
| --- | --- | --- | --- |
| Control Plan | CP-HIP-001 Rev A | CNC Femoral Stem - Machining Control Plan | In Review |
| Work Instruction | WI-CNC-001 Rev B | CNC Turning Setup & Operation - Ti-6Al-4V Stem | Released |
| Inspection Plan | IP-HIP-003 Rev A | CMM Dimensional Inspection - Femoral Stem | Released |
| IQ/OQ Protocol | IQ-COOL-001 | Coolant Interlock System - Installation Qualification | Pending |
| OQ/PQ Protocol | PQ-CNC-STEM-001 | Process Validation - CNC Machining Transfer | Planned |
| Risk Management | RM-HIP-2024-01 | Device Risk Management File (ISO 14971) | In Review |
| DFMEA | DFMEA-HIP-001 | Design FMEA - Cementless Femoral Stem | Released |
| SOP | SOP-QE-022 | pFMEA Creation and Maintenance Procedure | Released |

Archive Location:  [PLM System] -> Projects -> HIP-STEM-TRANSFER-2024 -> FMEA -> PFMEA-HIP-001_RevA.xlsx

AIAG-VDA FMEA 2019  |  21 CFR 820  |  ISO 13485

# 8. ISO 14971 Bridge

ISO 14971 RISK BRIDGE - best-effort mapping, QE review REQUIRED

| Failure Mode (pFMEA) | Hazard | Hazardous Situation | Harm | Severity (S) | P1 (Prob. of hazardous situation) | P2 (Prob. of harm) |
| --- | --- | --- | --- | --- | --- | --- |
| OD out of tolerance (oversize) | [derive] | [derive] | Implant-stem mismatch; potential intraoperative fracture or non-union - patient harm | 8 | [assess] | [assess] |
| Surface roughness Ra > 0.8 um on taper | [derive] | [derive] | Reduced osseointegration; long-term implant loosening - re-operation risk | 9 | [assess] | [assess] |
| Fixture slip / part shift mid-cycle | [derive] | [derive] | Geometric error cascade (runout, perpendicularity OOS); scrap or undetected escape | 7 | [assess] | [assess] |
| Coolant pressure drop / flow loss | [derive] | [derive] | Thermal damage to Ti substrate; metallurgical alteration; implant scrapped or escaped | 9 | [assess] | [assess] |
| Wrong program revision loaded (obsolete program) | [derive] | [derive] | Incorrect geometry machined; toolpath collision risk; patient dimension non-conformance | 8 | [assess] | [assess] |
| CMM inspection step skipped / bypassed | [derive] | [derive] | Non-conforming part released to sterilization and distribution - direct patient risk | 7 | [assess] | [assess] |

AIAG-VDA FMEA 2019  |  21 CFR 820  |  ISO 13485  |  ISO 14971:2019 - map P1xP2->Risk per device risk policy