## 1. Scope & Planning
STEP 1 - Scope & Planning ¦ CNC Machining - Ti-6Al-4V Femoral Stem
Field | Detail
Process Name | CNC Machining - Ti-6Al-4V Femoral Stem
Scope Start | Raw Ti-6Al-4V Bar Stock Receipt
Scope End | Final Dimensional Inspection / Release
Device Classification | Class III - PMA (21 CFR 814)
Standard Applied | AIAG-VDA FMEA 4th Edition 2019
Regulatory Framework | 21 CFR 820 (QSR) ¦ ISO 13485:2016
Device Description | Cementless Ti-6Al-4V Femoral Stem - Implantable Hip Prosthesis
Team Members | Quality Engineer, Process Engineer, Supplier Quality, R&D
FMEA Facilitator | [Your Name] - Quality Engineer
FMEA Number | PFMEA-HIP-001
Revision | Rev A
Date | 2026-06-22
Next Review Date | 6 months post-transfer validation
Status | Draft - Internal Review
AIAG-VDA FMEA 2019 ¦ 21 CFR 820 ¦ ISO 13485

## 2. Structure Analysis
STEP 2 - Structure Analysis ¦ System Hierarchy Breakdown
System Level | Element Name | Function Contribution
Level 1 - System | Manufacturing Cell | CNC Machining production environment
Level 2 - Subsystem | CNC Turning Station | Primary material removal and geometry creation
Level 3 - Component | Cutting Tool (Insert) | Executes material removal; controls Ra and geometry
Level 3 - Component | Workholding Fixture | Maintains part datum and prevents movement
Level 3 - Component | Coolant Delivery System | Controls chip evacuation and thermal stability
Level 3 - Component | CNC Part Program | Defines toolpath, feeds, speeds, and tolerances
Level 3 - Component | Operator / Setup Tech | Executes setup, monitors process, responds to alarms
Level 4 - Interface | CMM Inspection Station | In-process and final dimensional verification
Level 4 - Interface | Electronic Traveler (DHR) | Enforces operation sequence and records results
AIAG-VDA FMEA 2019 ¦ 21 CFR 820 ¦ ISO 13485

## 3. Function Analysis
STEP 3 - Function Analysis ¦ Intended Functions & Requirements
Process Element | Intended Function | Requirement Satisfied
Cutting Tool (Insert) | Remove material to achieve OD +/-0.010 mm and Ra <= 0.8 um on taper surface | Device drawing tolerance stack; osseointegration surface finish spec
Workholding Fixture | Maintain part concentricity <0.005 mm TIR during all turning operations | Geometric dimensioning: perpendicularity and runout callouts
Coolant Delivery System | Deliver coolant at >=50 PSI to maintain <60C tool-chip interface temperature | Ti-6Al-4V machining process spec; tool life and metallurgical integrity
CNC Part Program | Execute validated toolpath per ECO-approved revision; enforce feed/speed limits | Validated manufacturing process (IQ/OQ/PQ); 21 CFR 820.70(i) software control
Operator / Setup Tech | Load correct program revision, set datum offsets, perform first-article inspection | Work instruction WI-CNC-001; operator qualification records OQR-2024
CMM Inspection Station | Verify OD, ID, taper angle, and length within print tolerance post-machining | Inspection plan IP-HIP-003; acceptance criteria per device drawing Rev G
Electronic Traveler (DHR) | Enforce operation sequence gate checks; capture lot/SN, operator, and tool data | 21 CFR 820.184 Device History Record; ISO 13485 7.5.3 traceability
AIAG-VDA FMEA 2019 ¦ 21 CFR 820 ¦ ISO 13485

## 4. Failure Analysis
STEP 4 - Failure Analysis ¦ Failure Modes, Effects & Causes
Process Step | Failure Mode | Effect on Patient / Downstream Process | Severity (1-10) | Root Cause | Current Prevention Control | Current Detection Control
CNC Turning - OD | OD out of tolerance (oversize) | Implant-stem mismatch; potential intraoperative fracture or non-union - patient harm | 8 | Tool wear beyond limit; incorrect offset entry; thermal expansion of Ti | Insert change interval SOP; offset verification at setup | Post-op CMM inspection; SPC chart OD trend
CNC Turning - Surface | Surface roughness Ra > 0.8 um on taper | Reduced osseointegration; long-term implant loosening - re-operation risk | 9 | Worn or chipped insert; inadequate coolant flow; excessive feed rate | Profilometer check per lot; insert change per interval | Profilometer measurement per part; visual inspection
Workholding - Fixture | Fixture slip / part shift mid-cycle | Geometric error cascade (runout, perpendicularity OOS); scrap or undetected escape | 7 | Improper clamping torque; fixture wear; incorrect locating pin seating | Torque spec on setup sheet; periodic fixture calibration | First-article runout check; operator pre-run confirmation
Coolant Delivery | Coolant pressure drop / flow loss | Thermal damage to Ti substrate; metallurgical alteration; implant scrapped or escaped | 9 | Clogged filter; pump failure; hose leak; operator valve error | Coolant PM schedule; filter change interval | Visual observation only - no automated interlock (GAP)
CNC Program - Revision | Wrong program revision loaded (obsolete program) | Incorrect geometry machined; toolpath collision risk; patient dimension non-conformance | 8 | Paper traveler allows manual program selection; no enforced program lock | Program naming convention; operator training | Setup verification checklist; first-article measurement
Inspection - CMM Step | CMM inspection step skipped / bypassed | Non-conforming part released to sterilization and distribution - direct patient risk | 7 | Production schedule pressure; traveler signature forgery; inadequate gate control | Work instruction WI-CMM-002; supervisor sign-off | DHR review at final release - detection is end-of-line only (GAP)
AIAG-VDA FMEA 2019 ¦ 21 CFR 820 ¦ ISO 13485

## 5. Risk Rating
AIAG-VDA 2019 Action Priority (AP) -- NOT Traditional RPN
AIAG-VDA 2019 replaces Risk Priority Number (RPN = SxOxD) with Action Priority (AP: High / Medium / Low). Rationale for medical devices: a failure with Severity 9 that is NEVER detected (D=10) must be actioned immediately regardless of occurrence rate - RPN arithmetic can mask this by averaging scores. AP weights Detection gaps more heavily, reflecting that an undetected failure reaching a patient is catastrophic. The AP lookup table (not multiplication) drives priority.
AP Lookup: S9-10+D7-10->H ¦ S9-10+D4-6+O>=4->H ¦ S7-8+D7-10+O>=4->H ¦ All others->M or L
STEP 5 - Risk Rating ¦ AIAG-VDA Action Priority Table
Failure Mode | S (Severity 1-10) | O (Occurrence 1-10) | D (Detection 1-10) | Action Priority | AP Rationale | Legacy RPN (Reference Only)
OD out of tolerance (oversize) | 8 | 4 | 3 | L | S=7-8 + D<=3: detection adequate for severity level | 96
Surface roughness Ra > 0.8 um | 9 | 3 | 2 | M | S>=9 + D<=3: strong detection mitigates critical severity | 54
Fixture slip / part shift mid-cycle | 7 | 3 | 4 | L | S=7-8 + D=4-6 + O<4: controlled risk | 84
Coolant pressure drop / flow loss | 9 | 3 | 5 | M | S>=9 + D=4-6 + O<4: rare critical but detection acceptable | 135
Wrong program revision loaded | 8 | 2 | 3 | L | S=7-8 + D<=3: detection adequate for severity level | 48
CMM inspection step skipped | 7 | 4 | 7 | H | S=7-8 + D>=7 + O>=4: high-severity, frequent, poorly detected | 196
AIAG-VDA FMEA 2019 ¦ 21 CFR 820 ¦ ISO 13485

## 6. Optimization
STEP 6 - Optimization / Corrective Actions ¦ Risk Reduction Plan
Initial AP | Failure Mode | Action Description | Action Owner | Target Date | Revised S | Revised O | Revised D | Revised AP
H | Coolant pressure drop / flow loss | Install in-line coolant flow sensor with machine interlock: auto-feed hold + alarm if PSI < 45. Validate per IQ/OQ protocol IQ-COOL-001. Update PFMEA and Control Plan CP-HIP-001. | Process Engineer | 2026-Q3 | 9 | 2 | 2 | M
H | Wrong program revision loaded | Implement electronic traveler (eDHR) with barcode-locked program selection. Only approved ECO revision loads; prior revisions auto-archived. Validate per 21 CFR 820.70(i). Training completion rate 100% before go-live. | Quality Engineer + IT | 2026-Q4 | 8 | 2 | 1 | L
M | Surface roughness Ra > 0.8 um | Reduce insert change interval from 150 to 100 parts; implement profilometer check every 50 parts (sample basis). Update PM schedule and WI-CNC-001 Rev B. | Quality Engineer | 2026-Q3 | 9 | 2 | 1 | M
H | CMM inspection step skipped | Add mandatory CMM gate-check in eDHR workflow: part cannot advance until CMM pass record is uploaded. CMM data auto-linked to DHR lot record. Supervisor override requires Level III QE approval. | Quality Engineer | 2026-Q4 | 7 | 3 | 2 | L
AIAG-VDA FMEA 2019 ¦ 21 CFR 820 ¦ ISO 13485

## 7. Results Summary
STEP 7 - Results & Documentation Summary
pFMEA Metrics Summary (auto-computed)
Total Failure Modes Analyzed |  |  | 6
Initial High (H) Action Priority |  |  | 1
High AP After Actions Implemented |  |  | 0
Medium (M) AP - Initial |  |  | 2
Medium (M) AP - After Actions |  |  | 2
Low (L) AP - After Actions |  |  | 4
Total Corrective Actions Assigned |  |  | 4
Actions with Owner Assigned |  |  | 4
Actions with Target Date |  |  | 4
Linked Quality System Documents
Document Type | Document Number | Title / Description | Status
Control Plan | CP-HIP-001 Rev A | CNC Femoral Stem - Machining Control Plan | In Review
Work Instruction | WI-CNC-001 Rev B | CNC Turning Setup & Operation - Ti-6Al-4V Stem | Released
Inspection Plan | IP-HIP-003 Rev A | CMM Dimensional Inspection - Femoral Stem | Released
IQ/OQ Protocol | IQ-COOL-001 | Coolant Interlock System - Installation Qualification | Pending
OQ/PQ Protocol | PQ-CNC-STEM-001 | Process Validation - CNC Machining Transfer | Planned
Risk Management | RM-HIP-2024-01 | Device Risk Management File (ISO 14971) | In Review
DFMEA | DFMEA-HIP-001 | Design FMEA - Cementless Femoral Stem | Released
SOP | SOP-QE-022 | pFMEA Creation and Maintenance Procedure | Released
Archive Location: [PLM System] -> Projects -> HIP-STEM-TRANSFER-2024 -> FMEA -> PFMEA-HIP-001_RevA.xlsx
AIAG-VDA FMEA 2019 ¦ 21 CFR 820 ¦ ISO 13485

## 8. ISO 14971 Bridge
ISO 14971 RISK BRIDGE - best-effort mapping, QE review REQUIRED
Failure Mode (pFMEA) | Hazard | Hazardous Situation | Harm | Severity (S) | P1 (Prob. of hazardous situation) | P2 (Prob. of harm)
OD out of tolerance (oversize) | [derive] | [derive] | Implant-stem mismatch; potential intraoperative fracture or non-union - patient harm | 8 | [assess] | [assess]
Surface roughness Ra > 0.8 um on taper | [derive] | [derive] | Reduced osseointegration; long-term implant loosening - re-operation risk | 9 | [assess] | [assess]
Fixture slip / part shift mid-cycle | [derive] | [derive] | Geometric error cascade (runout, perpendicularity OOS); scrap or undetected escape | 7 | [assess] | [assess]
Coolant pressure drop / flow loss | [derive] | [derive] | Thermal damage to Ti substrate; metallurgical alteration; implant scrapped or escaped | 9 | [assess] | [assess]
Wrong program revision loaded (obsolete program) | [derive] | [derive] | Incorrect geometry machined; toolpath collision risk; patient dimension non-conformance | 8 | [assess] | [assess]
CMM inspection step skipped / bypassed | [derive] | [derive] | Non-conforming part released to sterilization and distribution - direct patient risk | 7 | [assess] | [assess]
AIAG-VDA FMEA 2019 ¦ 21 CFR 820 ¦ ISO 13485 ¦ ISO 14971:2019 - map P1xP2->Risk per device risk policy