# 1. Scope & Planning

STEP 1 - Scope & Planning  |  Injection Molding - PEEK Lumbar Interbody Fusion Cage

| Field | Detail |
| --- | --- |
| Process Name | Injection Molding - PEEK Lumbar Interbody Fusion Cage |
| Scope Start | Raw PEEK Pellet Receipt & Material Verification |
| Scope End | Final Dimensional Inspection & Lot Release |
| Device Classification | Class II - 510(k) (21 CFR 888.3070) |
| Standard Applied | AIAG-VDA FMEA 4th Edition 2019 |
| Regulatory Framework | 21 CFR 820 (QSR) ¦ ISO 13485:2016 ¦ ASTM F2026 |
| Device Description | PEEK Lumbar Interbody Fusion Cage - Spinal Implant |
| Team Members | Mold Engineer, Quality Engineer, Process Engineer, Supplier Quality, R&D |
| FMEA Facilitator | [Your Name] - Quality Engineer |
| FMEA Number | PFMEA-SPINE-001 |
| Revision | Rev A |
| Date | 2026-06-22 |
| Next Review Date | 6 months post-transfer validation |
| Status | Draft - Internal Review |

AIAG-VDA FMEA 2019  |  21 CFR 820  |  ISO 13485

# 2. Structure Analysis

STEP 2 - Structure Analysis  |  System Hierarchy Breakdown

| System Level | Element Name | Function Contribution |
| --- | --- | --- |
| Level 1 - System | Injection Molding Cell | Thermoplastic implant manufacturing environment |
| Level 2 - Subsystem | Injection Molding Press | Melts and injects PEEK at controlled temp/pressure |
| Level 3 - Component | Mold Tool (Cavity/Core) | Defines part geometry, surface finish, and tolerances |
| Level 3 - Component | Barrel / Screw Assembly | Plasticizes PEEK; controls melt temperature and shear |
| Level 3 - Component | Hot Runner / Gate System | Delivers melt uniformly; controls gate vestige height |
| Level 3 - Component | Material Handling & Drying | Ensures PEEK moisture <0.02% before processing |
| Level 3 - Component | Process Controller (HMI) | Enforces validated parameter recipe (IQ/OQ/PQ) |
| Level 4 - Interface | CMM / Vision Inspection Station | Dimensional and visual verification post-mold |
| Level 4 - Interface | Electronic Traveler (DHR) | Records lot, material CoA, operator, and process params |

AIAG-VDA FMEA 2019  |  21 CFR 820  |  ISO 13485

# 3. Function Analysis

STEP 3 - Function Analysis  |  Intended Functions & Requirements

| Process Element | Intended Function | Requirement Satisfied |
| --- | --- | --- |
| Mold Tool (Cavity/Core) | Form cage geometry to OD/ID and footprint tolerances per drawing; achieve Ra <= 1.6 um on endplates | Device drawing dimensional stack; osseointegration surface spec per ASTM F2026 |
| Barrel / Screw Assembly | Maintain PEEK melt temperature 380-400 C; prevent thermal degradation of polymer | PEEK processing spec; material property retention per ISO 10993-1 biocompatibility |
| Hot Runner / Gate System | Deliver bubble-free melt shot; gate vestige height <= 0.3 mm above parting line | Device drawing gate vestige callout; visual inspection acceptance criteria |
| Material Handling & Drying | Dry PEEK pellets to moisture <= 0.02% at 150 C for >= 4 hours before molding | PEEK processing guideline; prevention of hydrolytic degradation and splay defects |
| Process Controller (HMI) | Lock and enforce validated parameter recipe; flag deviation >5% from setpoint | Validated manufacturing process (IQ/OQ/PQ); 21 CFR 820.70(i) software control |
| CMM / Vision Inspection | Verify footprint, height, endplate geometry, and gate vestige against print tolerance | Inspection plan IP-SPINE-002; acceptance criteria per device drawing Rev D |
| Electronic Traveler (DHR) | Gate operation sequence; capture material lot, CoA link, operator, and process params | 21 CFR 820.184 Device History Record; ISO 13485 7.5.3 traceability |

AIAG-VDA FMEA 2019  |  21 CFR 820  |  ISO 13485

# 4. Failure Analysis

STEP 4 - Failure Analysis  |  Failure Modes, Effects & Causes

| Process Step | Failure Mode | Effect on Patient / Downstream Process | Severity (1-10) | Root Cause | Current Prevention Control | Current Detection Control |
| --- | --- | --- | --- | --- | --- | --- |
| Injection Molding - Parting Line | Flash / fin exceeding 0.1 mm at parting line | Sharp edge risk during implant handling; potential particulate in surgical field | 6 | Insufficient clamp force; worn parting line; mold damage | Clamp force setpoint in validated recipe; mold PM schedule | 100% visual inspection per WI-VIS-003; AQL sampling CMM |
| Injection Molding - Geometry | Outer footprint or height OOS (dimensional deviation) | Implant-endplate mismatch; subsidence risk; non-union - patient harm | 8 | Mold wear / thermal expansion; incorrect recipe offset; shot-to-shot variation | Recipe lock in HMI; mold qualification per OQ; SPC on critical dims | CMM inspection per sample plan SP-SPINE-01; first-article 100% |
| Material Handling | Wrong PEEK grade loaded (incorrect molecular weight / filler) | Incorrect mechanical properties; implant fracture or fatigue failure in vivo | 9 | Manual material selection; similar pellet appearance; label misread | Material labeling SOP; operator training | Visual label check only - no barcode/CoA lock (DETECTION GAP) |
| Injection Molding - Fill | Short shot / incomplete mold fill | Structural porosity in load-bearing cage; fracture risk under spinal loading | 7 | Insufficient injection pressure; blocked gate; material viscosity variation | Injection pressure setpoint and transfer position in validated recipe | Visual inspection - porosity may be subsurface and missed (GAP) |
| Gate / Trim | Gate vestige height > 0.3 mm above parting surface | Mechanical interference with endplate; surgeon re-work; particulate generation | 7 | Gate freeze inconsistency; dwell time variation; manual trim variability | Dwell time in recipe; gate vestige spec on drawing | First-article caliper check; periodic sampling only |
| Injection Molding - Surface | Endplate surface texture OOS (Ra > 1.6 um or peak pattern missing) | Reduced bone in-growth; long-term instability and re-operation risk | 8 | Worn or damaged mold texture; contamination on cavity surface; improper mold release | Mold surface inspection at PM; release agent SOP | Profilometer sample check per lot - not 100% (DETECTION GAP) |

AIAG-VDA FMEA 2019  |  21 CFR 820  |  ISO 13485

# 5. Risk Rating

AIAG-VDA 2019 Action Priority (AP) -- NOT Traditional RPN

AIAG-VDA 2019 replaces RPN (SxOxD) with Action Priority (AP: High / Medium / Low). For spinal implants, an undetected material substitution (wrong PEEK grade) reaching a patient represents a catastrophic fracture risk regardless of how rarely it occurs. AP weights Detection gaps heavily - a D=7 with any S>=9 is automatically High priority.

AP Lookup: S9-10+D7-10->H | S9-10+D4-6+O>=4->H | S7-8+D7-10+O>=4->H | All others->M or L

STEP 5 - Risk Rating  |  AIAG-VDA Action Priority Table

| Failure Mode | S (Severity 1-10) | O (Occurrence 1-10) | D (Detection 1-10) | Action Priority | AP Rationale | Legacy RPN (Reference Only) |
| --- | --- | --- | --- | --- | --- | --- |
| Flash / fin at parting line | 6 | 4 | 3 | L | Lower severity with adequate controls | 72 |
| Outer footprint or height OOS | 8 | 4 | 5 | M | S=7-8 + D=4-6 + O>=4: moderate risk profile | 160 |
| Wrong PEEK grade loaded | 9 | 3 | 7 | H | S>=9 + D>=7: undetected critical failure - patient risk | 189 |
| Short shot / incomplete fill | 7 | 3 | 2 | L | S=7-8 + D<=3: detection adequate for severity level | 42 |
| Gate vestige height > 0.3 mm | 7 | 4 | 4 | M | S=7-8 + D=4-6 + O>=4: moderate risk profile | 112 |
| Endplate surface texture OOS | 8 | 5 | 7 | H | S=7-8 + D>=7 + O>=4: high-severity, frequent, poorly detected | 280 |

AIAG-VDA FMEA 2019  |  21 CFR 820  |  ISO 13485

# 6. Optimization

STEP 6 - Optimization / Corrective Actions  |  Risk Reduction Plan

| Initial AP | Failure Mode | Action Description | Action Owner | Target Date | Revised S | Revised O | Revised D | Revised AP |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| H | Wrong PEEK grade loaded | Implement material barcode scan at press hopper: eDHR locks out start until scanned lot matches approved CoA on traveler. Certificate of Analysis auto-linked to DHR. Validate per 21 CFR 820.70(i); 100% operator training before go-live. | Quality Engineer | 2026-Q3 | 9 | 2 | 2 | M |
| H | Endplate surface texture OOS | Install inline profilometer at ejection station: 100% Ra measurement, SPC alert at Ra > 1.4 um. Mold surface inspection added to PM checklist at every 5,000 shots. Update Control Plan CP-SPINE-001 and WI-MOLD-002 Rev B. | Process Engineer | 2026-Q3 | 8 | 3 | 3 | L |
| M | Outer footprint or height OOS | Add 100% vision system check for critical footprint and height at ejection. Non-conforming parts automatically rejected to quarantine bin. SPC charts for cavity-specific dimensional trending. | Quality Engineer | 2026-Q4 | 8 | 2 | 2 | L |
| M | Gate vestige height > 0.3 mm | Add gate vestige height to mandatory first-article CMM program (currently omitted). Increase periodic sampling from 1/lot to 5/lot until Cpk >= 1.67 demonstrated over 30 lots. | Quality Engineer | 2026-Q4 | 7 | 3 | 2 | L |

AIAG-VDA FMEA 2019  |  21 CFR 820  |  ISO 13485

# 7. Results Summary

STEP 7 - Results & Documentation Summary

pFMEA Metrics Summary (auto-computed)

| Total Failure Modes Analyzed | 6 |
| --- | --- |
| Initial High (H) Action Priority | 2 |
| High AP After Actions Implemented | 0 |
| Medium (M) AP - Initial | 2 |
| Medium (M) AP - After Actions | 1 |
| Low (L) AP - After Actions | 5 |
| Total Corrective Actions Assigned | 4 |
| Actions with Owner Assigned | 4 |
| Actions with Target Date | 4 |

Linked Quality System Documents

| Document Type | Document Number | Title / Description | Status |
| --- | --- | --- | --- |
| Control Plan | CP-SPINE-001 Rev A | Injection Molding - PEEK Spinal Cage Control Plan | In Review |
| Work Instruction | WI-MOLD-002 Rev B | Injection Molding Setup & Operation - PEEK Cage | Released |
| Inspection Plan | IP-SPINE-002 Rev A | CMM & Vision Inspection - Lumbar Interbody Cage | Released |
| IQ/OQ Protocol | IQ-MOLD-003 | Injection Press Qualification - PEEK Processing Parameters | Pending |
| PQ Protocol | PQ-SPINE-CAGE-001 | Process Validation - PEEK Cage Injection Molding Transfer | Planned |
| Risk Management | RM-SPINE-2024-02 | Device Risk Management File (ISO 14971) - Fusion Cage | In Review |
| DFMEA | DFMEA-SPINE-001 | Design FMEA - PEEK Lumbar Interbody Fusion Cage | Released |
| SOP | SOP-QE-022 | pFMEA Creation and Maintenance Procedure | Released |

Archive Location:  [PLM System] -> Projects -> SPINE-CAGE-TRANSFER-2024 -> FMEA -> PFMEA-SPINE-001_RevA.xlsx

AIAG-VDA FMEA 2019  |  21 CFR 820  |  ISO 13485