## 1. Scope & Planning
STEP 1 - Scope & Planning ¦ Injection Molding - PEEK Lumbar Interbody Fusion Cage
Field | Detail
Process Name | Injection Molding - PEEK Lumbar Interbody Fusion Cage
Scope Start | Raw PEEK Pellet Receipt & Material Verification
Scope End | Final Dimensional Inspection & Lot Release
Device Classification | Class II - 510(k) (21 CFR 888.3070)
Standard Applied | AIAG-VDA FMEA 4th Edition 2019
Regulatory Framework | 21 CFR 820 (QSR) ¦ ISO 13485:2016 ¦ ASTM F2026
Device Description | PEEK Lumbar Interbody Fusion Cage - Spinal Implant
Team Members | Mold Engineer, Quality Engineer, Process Engineer, Supplier Quality, R&D
FMEA Facilitator | [Your Name] - Quality Engineer
FMEA Number | PFMEA-SPINE-001
Revision | Rev A
Date | 2026-06-22
Next Review Date | 6 months post-transfer validation
Status | Draft - Internal Review
AIAG-VDA FMEA 2019 ¦ 21 CFR 820 ¦ ISO 13485

## 2. Structure Analysis
STEP 2 - Structure Analysis ¦ System Hierarchy Breakdown
System Level | Element Name | Function Contribution
Level 1 - System | Injection Molding Cell | Thermoplastic implant manufacturing environment
Level 2 - Subsystem | Injection Molding Press | Melts and injects PEEK at controlled temp/pressure
Level 3 - Component | Mold Tool (Cavity/Core) | Defines part geometry, surface finish, and tolerances
Level 3 - Component | Barrel / Screw Assembly | Plasticizes PEEK; controls melt temperature and shear
Level 3 - Component | Hot Runner / Gate System | Delivers melt uniformly; controls gate vestige height
Level 3 - Component | Material Handling & Drying | Ensures PEEK moisture <0.02% before processing
Level 3 - Component | Process Controller (HMI) | Enforces validated parameter recipe (IQ/OQ/PQ)
Level 4 - Interface | CMM / Vision Inspection Station | Dimensional and visual verification post-mold
Level 4 - Interface | Electronic Traveler (DHR) | Records lot, material CoA, operator, and process params
AIAG-VDA FMEA 2019 ¦ 21 CFR 820 ¦ ISO 13485

## 3. Function Analysis
STEP 3 - Function Analysis ¦ Intended Functions & Requirements
Process Element | Intended Function | Requirement Satisfied
Mold Tool (Cavity/Core) | Form cage geometry to OD/ID and footprint tolerances per drawing; achieve Ra <= 1.6 um on endplates | Device drawing dimensional stack; osseointegration surface spec per ASTM F2026
Barrel / Screw Assembly | Maintain PEEK melt temperature 380-400 C; prevent thermal degradation of polymer | PEEK processing spec; material property retention per ISO 10993-1 biocompatibility
Hot Runner / Gate System | Deliver bubble-free melt shot; gate vestige height <= 0.3 mm above parting line | Device drawing gate vestige callout; visual inspection acceptance criteria
Material Handling & Drying | Dry PEEK pellets to moisture <= 0.02% at 150 C for >= 4 hours before molding | PEEK processing guideline; prevention of hydrolytic degradation and splay defects
Process Controller (HMI) | Lock and enforce validated parameter recipe; flag deviation >5% from setpoint | Validated manufacturing process (IQ/OQ/PQ); 21 CFR 820.70(i) software control
CMM / Vision Inspection | Verify footprint, height, endplate geometry, and gate vestige against print tolerance | Inspection plan IP-SPINE-002; acceptance criteria per device drawing Rev D
Electronic Traveler (DHR) | Gate operation sequence; capture material lot, CoA link, operator, and process params | 21 CFR 820.184 Device History Record; ISO 13485 7.5.3 traceability
AIAG-VDA FMEA 2019 ¦ 21 CFR 820 ¦ ISO 13485

## 4. Failure Analysis
STEP 4 - Failure Analysis ¦ Failure Modes, Effects & Causes
Process Step | Failure Mode | Effect on Patient / Downstream Process | Severity (1-10) | Root Cause | Current Prevention Control | Current Detection Control
Injection Molding - Parting Line | Flash / fin exceeding 0.1 mm at parting line | Sharp edge risk during implant handling; potential particulate in surgical field | 6 | Insufficient clamp force; worn parting line; mold damage | Clamp force setpoint in validated recipe; mold PM schedule | 100% visual inspection per WI-VIS-003; AQL sampling CMM
Injection Molding - Geometry | Outer footprint or height OOS (dimensional deviation) | Implant-endplate mismatch; subsidence risk; non-union - patient harm | 8 | Mold wear / thermal expansion; incorrect recipe offset; shot-to-shot variation | Recipe lock in HMI; mold qualification per OQ; SPC on critical dims | CMM inspection per sample plan SP-SPINE-01; first-article 100%
Material Handling | Wrong PEEK grade loaded (incorrect molecular weight / filler) | Incorrect mechanical properties; implant fracture or fatigue failure in vivo | 9 | Manual material selection; similar pellet appearance; label misread | Material labeling SOP; operator training | Visual label check only - no barcode/CoA lock (DETECTION GAP)
Injection Molding - Fill | Short shot / incomplete mold fill | Structural porosity in load-bearing cage; fracture risk under spinal loading | 7 | Insufficient injection pressure; blocked gate; material viscosity variation | Injection pressure setpoint and transfer position in validated recipe | Visual inspection - porosity may be subsurface and missed (GAP)
Gate / Trim | Gate vestige height > 0.3 mm above parting surface | Mechanical interference with endplate; surgeon re-work; particulate generation | 7 | Gate freeze inconsistency; dwell time variation; manual trim variability | Dwell time in recipe; gate vestige spec on drawing | First-article caliper check; periodic sampling only
Injection Molding - Surface | Endplate surface texture OOS (Ra > 1.6 um or peak pattern missing) | Reduced bone in-growth; long-term instability and re-operation risk | 8 | Worn or damaged mold texture; contamination on cavity surface; improper mold release | Mold surface inspection at PM; release agent SOP | Profilometer sample check per lot - not 100% (DETECTION GAP)
AIAG-VDA FMEA 2019 ¦ 21 CFR 820 ¦ ISO 13485

## 5. Risk Rating
AIAG-VDA 2019 Action Priority (AP) -- NOT Traditional RPN
AIAG-VDA 2019 replaces RPN (SxOxD) with Action Priority (AP: High / Medium / Low). For spinal implants, an undetected material substitution (wrong PEEK grade) reaching a patient represents a catastrophic fracture risk regardless of how rarely it occurs. AP weights Detection gaps heavily - a D=7 with any S>=9 is automatically High priority.
AP Lookup: S9-10+D7-10->H ¦ S9-10+D4-6+O>=4->H ¦ S7-8+D7-10+O>=4->H ¦ All others->M or L
STEP 5 - Risk Rating ¦ AIAG-VDA Action Priority Table
Failure Mode | S (Severity 1-10) | O (Occurrence 1-10) | D (Detection 1-10) | Action Priority | AP Rationale | Legacy RPN (Reference Only)
Flash / fin at parting line | 6 | 4 | 3 | L | Lower severity with adequate controls | 72
Outer footprint or height OOS | 8 | 4 | 5 | M | S=7-8 + D=4-6 + O>=4: moderate risk profile | 160
Wrong PEEK grade loaded | 9 | 3 | 7 | H | S>=9 + D>=7: undetected critical failure - patient risk | 189
Short shot / incomplete fill | 7 | 3 | 2 | L | S=7-8 + D<=3: detection adequate for severity level | 42
Gate vestige height > 0.3 mm | 7 | 4 | 4 | M | S=7-8 + D=4-6 + O>=4: moderate risk profile | 112
Endplate surface texture OOS | 8 | 5 | 7 | H | S=7-8 + D>=7 + O>=4: high-severity, frequent, poorly detected | 280
AIAG-VDA FMEA 2019 ¦ 21 CFR 820 ¦ ISO 13485

## 6. Optimization
STEP 6 - Optimization / Corrective Actions ¦ Risk Reduction Plan
Initial AP | Failure Mode | Action Description | Action Owner | Target Date | Revised S | Revised O | Revised D | Revised AP
H | Wrong PEEK grade loaded | Implement material barcode scan at press hopper: eDHR locks out start until scanned lot matches approved CoA on traveler. Certificate of Analysis auto-linked to DHR. Validate per 21 CFR 820.70(i); 100% operator training before go-live. | Quality Engineer | 2026-Q3 | 9 | 2 | 2 | M
H | Endplate surface texture OOS | Install inline profilometer at ejection station: 100% Ra measurement, SPC alert at Ra > 1.4 um. Mold surface inspection added to PM checklist at every 5,000 shots. Update Control Plan CP-SPINE-001 and WI-MOLD-002 Rev B. | Process Engineer | 2026-Q3 | 8 | 3 | 3 | L
M | Outer footprint or height OOS | Add 100% vision system check for critical footprint and height at ejection. Non-conforming parts automatically rejected to quarantine bin. SPC charts for cavity-specific dimensional trending. | Quality Engineer | 2026-Q4 | 8 | 2 | 2 | L
M | Gate vestige height > 0.3 mm | Add gate vestige height to mandatory first-article CMM program (currently omitted). Increase periodic sampling from 1/lot to 5/lot until Cpk >= 1.67 demonstrated over 30 lots. | Quality Engineer | 2026-Q4 | 7 | 3 | 2 | L
AIAG-VDA FMEA 2019 ¦ 21 CFR 820 ¦ ISO 13485

## 7. Results Summary
STEP 7 - Results & Documentation Summary
pFMEA Metrics Summary (auto-computed)
Total Failure Modes Analyzed |  |  | 6
Initial High (H) Action Priority |  |  | 2
High AP After Actions Implemented |  |  | 0
Medium (M) AP - Initial |  |  | 2
Medium (M) AP - After Actions |  |  | 1
Low (L) AP - After Actions |  |  | 5
Total Corrective Actions Assigned |  |  | 4
Actions with Owner Assigned |  |  | 4
Actions with Target Date |  |  | 4
Linked Quality System Documents
Document Type | Document Number | Title / Description | Status
Control Plan | CP-SPINE-001 Rev A | Injection Molding - PEEK Spinal Cage Control Plan | In Review
Work Instruction | WI-MOLD-002 Rev B | Injection Molding Setup & Operation - PEEK Cage | Released
Inspection Plan | IP-SPINE-002 Rev A | CMM & Vision Inspection - Lumbar Interbody Cage | Released
IQ/OQ Protocol | IQ-MOLD-003 | Injection Press Qualification - PEEK Processing Parameters | Pending
PQ Protocol | PQ-SPINE-CAGE-001 | Process Validation - PEEK Cage Injection Molding Transfer | Planned
Risk Management | RM-SPINE-2024-02 | Device Risk Management File (ISO 14971) - Fusion Cage | In Review
DFMEA | DFMEA-SPINE-001 | Design FMEA - PEEK Lumbar Interbody Fusion Cage | Released
SOP | SOP-QE-022 | pFMEA Creation and Maintenance Procedure | Released
Archive Location: [PLM System] -> Projects -> SPINE-CAGE-TRANSFER-2024 -> FMEA -> PFMEA-SPINE-001_RevA.xlsx
AIAG-VDA FMEA 2019 ¦ 21 CFR 820 ¦ ISO 13485