## 1. Scope & Planning
STEP 1 - Scope & Planning ¦ Sterile Barrier Packaging Seal - Tyvek-Film Pouch (Class III Implants)
Field | Detail
Process Name | Sterile Barrier Packaging Seal - Tyvek-Film Pouch (Class III Implants)
Scope Start | Pouching / Loading of Sterilized Implant
Scope End | Sealing, Inspection, Labeling & Distribution Release
Device Classification | Class III - PMA support process (21 CFR 814)
Standard Applied | AIAG-VDA FMEA 4th Edition 2019
Regulatory Framework | 21 CFR 820 (QSR) ¦ ISO 13485:2016 ¦ ISO 11607-1:2019
Device Description | Sterile Tyvek-Film Pouch - Primary Sterile Barrier for Class III Implants
Team Members | Packaging Engineer, Quality Engineer, Sterilization Specialist, Supplier Quality
FMEA Facilitator | [Your Name] - Quality Engineer
FMEA Number | PFMEA-PKG-001
Revision | Rev A
Date | 2026-06-22
Next Review Date | Annual review per SOP-QE-022 or upon process change
Status | Draft - Internal Review
AIAG-VDA FMEA 2019 ¦ 21 CFR 820 ¦ ISO 13485 ¦ ISO 11607

## 2. Structure Analysis
STEP 2 - Structure Analysis ¦ System Hierarchy Breakdown
System Level | Element Name | Function Contribution
Level 1 - System | Sterile Packaging Line | End-to-end sterile barrier formation and validation
Level 2 - Subsystem | Rotary Heat Sealer | Forms hermetic seal on three open edges of Tyvek-film pouch
Level 3 - Component | Sealing Jaw / Platen | Applies controlled heat and pressure to create seal bond
Level 3 - Component | Temperature Controller (PID) | Maintains seal jaw within validated temperature window
Level 3 - Component | Dwell Timer | Controls contact time for complete seal formation
Level 3 - Component | Pouch Stock (Tyvek + Film) | Provides sterile barrier material with validated bond surface
Level 3 - Component | Operator / Packaging Technician | Loads implant, positions pouch, initiates seal cycle
Level 4 - Interface | Seal Integrity Test Station | Bubble emission (ASTM F2096) and peel strength verification
Level 4 - Interface | Label Applicator / eDHR | Applies UDI label; records lot, operator, and seal params
AIAG-VDA FMEA 2019 ¦ 21 CFR 820 ¦ ISO 13485 ¦ ISO 11607

## 3. Function Analysis
STEP 3 - Function Analysis ¦ Intended Functions & Requirements
Process Element | Intended Function | Requirement Satisfied
Rotary Heat Sealer | Form a continuous, channel-free hermetic seal on three sides; seal width >= 6 mm per ISO 11607-1 | ISO 11607-1 Annex A: seal width and channel-defect acceptance criteria
Temperature Controller (PID) | Maintain sealing jaw at 175 +/-5 C (validated window per OQ); alarm on deviation > 3 C | Seal OQ temperature mapping; process validation PQ-PKG-SEAL-001
Dwell Timer | Hold seal pressure for 1.2 +/-0.1 s (validated); prevent under-dwell causing weak seal | Seal strength OQ data; minimum peel force >= 1.5 N/15 mm per ASTM F88
Pouch Stock (Tyvek + Film) | Maintain microbial barrier and seal-ability; stored per supplier IFU (humidity/temp controlled) | ISO 11607-1 material qualification; supplier CoA per receiving SOP
Operator / Packaging Tech | Load single implant per pouch, orient correctly, position seal edge within fixture guide | Work instruction WI-PKG-001; operator qualification and annual re-certification
Seal Integrity Test Station | Detect seal channels >1 mm via bubble emission (ASTM F2096); measure peel strength per ASTM F88 | ISO 11607-1 Annex A2.2; AQL sampling plan SP-PKG-02
Label Applicator / eDHR | Apply correct UDI label with lot/SN; capture seal temp, dwell, operator ID in DHR | 21 CFR 830 UDI; 21 CFR 820.184 DHR; ISO 13485 7.5.3 traceability
AIAG-VDA FMEA 2019 ¦ 21 CFR 820 ¦ ISO 13485 ¦ ISO 11607

## 4. Failure Analysis
STEP 4 - Failure Analysis ¦ Failure Modes, Effects & Causes
Process Step | Failure Mode | Effect on Patient / Downstream Process | Severity (1-10) | Root Cause | Current Prevention Control | Current Detection Control
Heat Sealing | Incomplete / channel seal defect (> 1 mm channel) | Loss of sterile barrier; microbial ingress; patient infection from non-sterile implant | 9 | Jaw misalignment; pouch wrinkle in seal zone; temperature excursion; worn platen | Seal jaw alignment verification at setup; temperature setpoint validated per OQ | Visual inspection per WI-VIS-004 - channels < 1 mm not reliably detected visually (GAP)
Seal Integrity | Pinhole or micro-leak in film layer (microbial ingress path) | Sterility breach; patient infection; implant recall and field safety corrective action | 10 | Film puncture from implant edge; handling damage; raw material defect | Pouch edge guard; gentle handling SOP; incoming film inspection | Bubble emission test samples only (ASTM F2096) - not 100% per lot (DETECTION GAP)
Temperature Control | Seal temperature out of validated range (< 170 C or > 180 C) | Under-seal: weak bond, field delamination risk. Over-seal: Tyvek fiber damage, barrier compromise | 8 | PID controller drift; thermocouple degradation; power supply fluctuation | PID calibration per PM schedule; operator pre-shift temperature verification | Process monitoring via controller display - no independent thermocouple alarm (GAP)
Dwell Time | Incorrect dwell time (operator override or timer fault) | Under-dwell: weak seal, peel strength below 1.5 N/15 mm; sterile barrier at risk in distribution | 8 | Operator bypasses timer; controller fault; recipe not locked | Timer setpoint in validated recipe; operator training | Peel strength sampled 5 pouches/lot per ASTM F88 - sub-lot variation not captured
Labeling | UDI label / content mismatch (wrong lot or implant size on label) | Incorrect device implanted; patient harm; regulatory non-conformance (21 CFR 830) | 8 | Manual label selection; look-alike label stock; no barcode verification at point of use | Label control SOP; label segregation; 2-person verification | Final QA label review before release - human verification only (GAP)
Post-Seal Handling | Package breach / seal peel during EtO sterilization or distribution transit | Sterility loss at point of use; patient infection; product recall | 9 | Over-stacking in sterilization load; inadequate secondary packaging cushioning | Stacking height limit in sterilization SOP; secondary packaging spec | Post-EtO visual inspection (sampling only); distribution simulation testing per ASTM D4169 annual
AIAG-VDA FMEA 2019 ¦ 21 CFR 820 ¦ ISO 13485 ¦ ISO 11607

## 5. Risk Rating
AIAG-VDA 2019 Action Priority (AP) -- NOT Traditional RPN
AIAG-VDA 2019 replaces RPN (SxOxD) with Action Priority (AP: High / Medium / Low). For sterile packaging, the consequence of an undetected seal defect is a non-sterile implant reaching the patient - a catastrophic outcome regardless of defect frequency. ISO 11607 requires validated sterile barrier integrity; AP prioritizes detection gaps accordingly. Any Severity >= 9 failure with Detection >= 7 is automatically High priority - RPN arithmetic can mask this by averaging a low Occurrence score.
AP Lookup: S9-10+D7-10->H ¦ S9-10+D4-6+O>=4->H ¦ S7-8+D7-10+O>=4->H ¦ All others->M or L
STEP 5 - Risk Rating ¦ AIAG-VDA Action Priority Table
Failure Mode | S (Severity 1-10) | O (Occurrence 1-10) | D (Detection 1-10) | Action Priority | AP Rationale | Legacy RPN (Reference Only)
Incomplete / channel seal defect | 9 | 4 | 6 | H | S>=9 + D=4-6 + O>=4: frequent critical w/ weak detection | 216
Pinhole / micro-leak in film layer | 10 | 2 | 8 | H | S>=9 + D>=7: undetected critical failure - patient risk | 160
Seal temperature out of range | 8 | 4 | 5 | M | S=7-8 + D=4-6 + O>=4: moderate risk profile | 160
Incorrect dwell time | 8 | 3 | 3 | L | S=7-8 + D<=3: detection adequate for severity level | 72
UDI label / content mismatch | 8 | 2 | 4 | L | S=7-8 + D=4-6 + O<4: controlled risk | 64
Package breach during EtO / transit | 9 | 3 | 5 | M | S>=9 + D=4-6 + O<4: rare critical but detection acceptable | 135
AIAG-VDA FMEA 2019 ¦ 21 CFR 820 ¦ ISO 13485 ¦ ISO 11607

## 6. Optimization
STEP 6 - Optimization / Corrective Actions ¦ Risk Reduction Plan
Initial AP | Failure Mode | Action Description | Action Owner | Target Date | Revised S | Revised O | Revised D | Revised AP
H | Incomplete / channel seal defect | Implement 100% bubble emission test per ASTM F2096 (replacing sampling). Add automated peel strength tester with SPC control chart - alert at < 1.8 N/15 mm. Update Control Plan CP-PKG-001 and sampling plan SP-PKG-02. Validate test equipment per IQ/OQ IQ-SIT-001. | Packaging Engineer + QE | 2026-Q3 | 9 | 2 | 3 | M
H | Pinhole / micro-leak in film layer | Add dye penetration test per ISO 11607-1 Annex A2.2 to every batch AQL (n=32, AQL 0.65). Implement incoming film inspection with pin-hole detector per incoming SOP. Add implant edge guard (foam insert) to pouching fixture. | Quality Engineer | 2026-Q3 | 10 | 2 | 4 | M
M | Seal temperature out of range | Install independent thermocouple with dual-channel alarm: feed-hold if jaw temp deviates > 3 C from setpoint. Validate alarm response per OQ protocol OQ-SEAL-002. Reduce PID calibration interval from 12 to 6 months. | Process Engineer | 2026-Q4 | 8 | 2 | 3 | L
M | Package breach during EtO / transit | Add post-EtO visual inspection of all pouches before labeling (100%, not sampling). Implement ASTM D4169 distribution simulation testing semi-annually (current: annual). Add foam separator between pouch layers in sterilization cassette. | Packaging + Sterilization PE | 2026-Q4 | 9 | 2 | 3 | M
AIAG-VDA FMEA 2019 ¦ 21 CFR 820 ¦ ISO 13485 ¦ ISO 11607

## 7. Results Summary
STEP 7 - Results & Documentation Summary
pFMEA Metrics Summary (auto-computed)
Total Failure Modes Analyzed |  |  | 6
Initial High (H) Action Priority |  |  | 2
High AP After Actions Implemented |  |  | 0
Medium (M) AP - Initial |  |  | 2
Medium (M) AP - After Actions |  |  | 3
Low (L) AP - After Actions |  |  | 3
Total Corrective Actions Assigned |  |  | 4
Actions with Owner Assigned |  |  | 4
Actions with Target Date |  |  | 4
Linked Quality System Documents
Document Type | Document Number | Title / Description | Status
Control Plan | CP-PKG-001 Rev A | Sterile Packaging - Heat Seal Control Plan | In Review
Work Instruction | WI-PKG-001 Rev C | Tyvek-Film Pouch Sealing - Setup & Operation | Released
Inspection Plan | SP-PKG-02 Rev B | Seal Integrity Sampling Plan - Bubble Emission & Peel Strength | Released
OQ Protocol | OQ-SEAL-002 | Heat Sealer Temperature & Dwell Time - Operational Qualification | In Progress
PQ Protocol | PQ-PKG-SEAL-001 | Process Validation - Sterile Packaging Transfer | Planned
Risk Management | RM-PKG-2024-01 | Packaging Risk Management File (ISO 14971) | In Review
Sterilization PQ | PQ-ETO-HIP-001 | EtO Sterilization Validation - Class III Implant Family | Released
SOP | SOP-QE-022 | pFMEA Creation and Maintenance Procedure | Released
Archive Location: [PLM System] -> Projects -> PKG-SEAL-TRANSFER-2024 -> FMEA -> PFMEA-PKG-001_RevA.xlsx
AIAG-VDA FMEA 2019 ¦ 21 CFR 820 ¦ ISO 13485 ¦ ISO 11607